care.

%

^f

4

Thescannercanbeoperatedinmanyways:

•initsleathercaseorflightcase.

•initsdockingcradleonatrolleyordesk

•initswallmountedcradle

To	print	results,	the	scanner	is	placed	in	its	cradle	and	if	required	can	stay	there

during	use.	More	conveniently	the	results	can	be	sent	via	its	USB	port	to	any

suitabledevice.ThePalmBladderScannerisahandhelddeviceandtheperfect

answer	for	easy	transportation	between	different	hospital	locations	or	for	use

in	primary	care	environment	such	as	a	continence	service	delivering	home

Technical Specification for Palm Bladder Scan

Scan method: 3D	Mechanical	sector	scanning

Probe freq.: 2.6	MHz	central	freq.

Scan depth: 180mm

Swing angle: 120	degree

Rotating angle: 180	degree

Screen size: 7	inches	TFT	color	LCD

Feature/Functions: Real-time	pre-scan	function	for	bladder	location	/

Sound	alarm	function	/

Bladder	image	and	volume	display	/	Scanned	image	storage	/	Software	upgrade	through	USB	/	Printer

build	in	cradle	/	Battery	consumption	management

Max.volume 	Measuring: 999ml

Difference: ±20%	(Volume	150ml)

±15%	(Volume	>	150ml)

Working environment: Temp.	+5	~	+40	Humidity	≤	80%

Atmosphere pressure: 700hpa-1060hpa

Power supply: Rechargable	lithium	aggregation	battery

Battery capacity: 6300mAh

Charging method: AC	Adapter,	mobile	charger

Input for AC adapter:

220V±10%	50/60Hz

Instrument dimension: 32cm	x	17cm	x	5cm

Weight: 2kg

Accessories: Trolley	or	wall-docking	station	Aluminium	look	case	or

leather	case	for	PBS	Delivered	With	batteries	(2x)	and	charger